Model 4000

X-Channel Extracellular Differential AC Amplifier

#### Labview Source Code

#### Version 0.1

# Table of Contents

Table of Contents	1

Model 4000 Labview  VI’s	2

SendReceive.vi   	2

WriteChannel.vi   	2

MantExp2str.vi   	2

Main.vi	2

Instrument Messaging	3

STANDARD INSTURMENT OVERVIEW	3

CUSTOMIZED INSTRUMENTS	4

MESSAGES	4

AMS_RESPONSE_ERROR	4

AMS_SLAVE_RESONSE	5

AMS_MSG_READ_FIRMWARE_VERSION	5

AMS_MSG_READ_HARDWARE_CONFIG	6

AMS_MSG_READ_INSTRUMENT_NAME	6

AMS_MSG_READ_BOX_AMOUNT	6

AMS_MSG_READ_SERIAL_NUMBER	6

AMS_MSG_READ_FLASH_BLOCK	8

AMS_MSG_LOAD_FLASH_BLOCK	8

AMS_MSG_SAVE_FLASH_BLOCK	9

AMS_MSG_WRITE_CHANNEL_VALUES	10

{hardware config block}	11

{channel config block}	11

{config value}	12

Revision History	13

# Model 4000 Labview  VI’s

Labview VI’s are set around the Main VI, and rely on subVI’s in the SubVI folder and front panel controls, or variables in the Controls folder. The VI’s provide communication to the instrument. .   Instrument communication is simplified by the VI’s, and is explained in detail in the Instrument Messaging section.  Communication is done over the a virtual serial port created by the usb device in the instrument.  The usb drivers need to be installed properly for the class to be used.  Current drivers can be downloaded from FTDI’s websites:

http://www.ftdichip.com/FTDrivers.htm.

## SendReceive.vi

This VI takes care of the basic instrument communication, and relies of Labview’s VISA resource for serial ports.  The Ring Input contains the command to execute on the instrument. SendBytes in a byte array with any data to send.  All messages get terminated automatically with the hexadecimal byte 7F.  VISA Write will send the data and depending on the command different lengths of data will be read with VISA Read and saved into a byte Array.  The returned byte array will be converted into a string to be returned to the calling VI along with the original byte array and the byte array length.

## WriteChannel.vi

This VI will take the channel cluster values along with the line frequency and reference bus and send these values to the instrument for immediate update of the instrument settings.

## MantExp2str.vi

This VI will take the instruments saved configuration settings that are in a mantissa and exponent format and convert it into a string.

## Main.vi

This VI allows setting one of the instrument settings, or all 32 at once.  Each time a setting is made without errors the Success light will light up.

# Instrument Messaging

## STANDARD INSTURMENT OVERVIEW

The Model 4000 MultiRecord consist of  32 independently controllable signal channels.  Multiple models can be cascaded together and controlled by one model, up to 256 channels.  For each channel, users can select via a  PC to USB to UART:

Enable/disable the channel

1 of 8 high pass filter settings

Enable/disable a notch filter

If the notch filter frequency is at 50 or 60Hz.

1 of 8 low pass filter settings

1 of 8 gain settings.

Additionally the users can set or read the instrument name, and read the revision and serial number.

To set any settings, an ASCII message is sent from the PC to firmware in the following format:

MessageVerb  MessageBytes (if applicable) MessageTerminator

The firmware will respond with:

StartMessage MessageNumber  Response_Verb Response_bytes EndMessage

The details of these messages is given the available message section.

For example the message for the PC to Read the Instrument Name would be in hexidecimal:

A6 7F

Where A6 is the MessageVerb, and 7F is the MessageTerminator, No Message bytes were necessary for this command  The Firmware would respond:

81 01 A7 4D 75 6C 74 69 2D 52 65 63 6F 72 64 20 41 6D 70 2E 00 81

Where 81 is the StartMessage indicator and the EndMessage indicator,

01 is the MessageNumber,

A7 is the MessageVerb,

4D 75 6C 74 69 2D 52 65 63 6F 72 64 20 41 6D 70 2E 00 is:

“Multi-Record Amp.<NUL>” (ASCII converted)

(REV 3 NOTE: there is no need for the ASCII format in this instrument –PC communication can be however we want it. The rest of the commands are really binary anyway ---)

All numerical settings for Write Channel Values are sent in ASCII format, so a 0 would be 30hex. All settings except for channel number have values less than 8, so only one byte is used for these variables.  Channel number may go up to 255, so two bytes are used for this setting: 00 to FF as the hexadecimal representation of 0 to 255. the high nibble goes first in ASCII format or 30h to 39h (0-9) and 41h to 46h(A-F), then the low nibble in ASCII format.

An example of this is the AMS_MSG_WRITE_CHANNEL_VALUES message:

If we want to set channel 47 on with the highpass filter to 100Hz, LineFrequency=60Hz,Notch=off,Reference=GND,low pass filter = 1kHz, and gain=50, then the following message would be sent:

B5 32 46 30 35 30 30 30 33 35 7F

Where B5 is the MessageVerb, and 32 46 is the hex value 2F, or 47 decimal, 30 35 30 30 30 33 35 are the filter and gain settings, and 7F is the MessageTerminator. The firmware would respond with the message:

81 02 C5 32 46 30 35 30 30 30 33 35 81

Where 81 is the StartMessage indicator and the EndMessage indicator,

02 is the MessageNumber,

C5 is the MessageVerb,

32 46 30 35 30 30 30 33 35 is:

“2 F 0 5 0 0 0 3 5” (ASCII converted)

## CUSTOMIZED INSTRUMENTS

A-M Systems, Inc. offers customers the option of ordering a factory-customized instrument with Low Pass and High Pass filter options and Gain options of their choice.  Additionally, these settings need not be common across all channels in the customized instrument.  To accommodate effective remote control, each instrument is able to communicate its factory preset hardware configuration, allowing the client application to adjust its user interface to accurately represent the available options implemented in the hardware.  The client application uses the AMS_MSG_READ_HARDWARE_CONFIG message to obtain this information from the instrument.  Details of this message format can be located in the section  titled Available Messages.

## MESSAGES

All communication the Model 4000 uses Hex Coded communication.  The communication includes a message verb defining the action to be performed, and any additional data needed for execution of the message.  Responses from the instrument are delivered to the client application beginning with a matching response verb to indicate the action performed, followed by any additional data sent with the response message.

Note:  Byte sequences enclosed in braces such as {hardware config bloc} are defined in detail at the end of this section.  Unless otherwise noted, all strings are variable in length and consist of up to 18 ASCII characters (1 byte each) followed by a required NULL terminator.

## AMS_RESPONSE_ERROR

If the Device receives a command it does not know it will send a Response as follows (again preceded by the StartMessage indicator, the MessageNumber and followed by the MessageTerminator)

Error Response Format (1 byte)	Length

Byte 0:	AMS_RESPONSE_ERROR (0xCD)	1

## AMS_SLAVE_RESONSE

If the is in slave mode it will only respond with this message.

Error Response Format (1 byte)	Length

Byte 0:	AMS_RESPONSE_ERROR (0xCE)	1

## AMS_MSG_READ_FIRMWARE_VERSION

Requests the build numbers for the internal processor firmware and the front panel LCD touch screen display firmware.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_READ_FIRMWARE_VERSION (0xA4)	1

Response Format (14 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_FIRMWARE_VERSION (0xA5)	1

Byte 1-13:	Firmware String (“YYYYMMDDHHMM”)	13

## AMS_MSG_READ_HARDWARE_CONFIG

Requests the hardware configuration data, which fully describes any hardware customizations present in the instrument.

Message Format (1 byte)	Length

Byte 0:	AMS_MSG_READ_HARDWARE_CONFIG (0xAA)	1

Response Format (321 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_HARDWARE_CONFIG (0xAB)	1

Byte 1-319:	{hardware config block}	320

## AMS_MSG_READ_INSTRUMENT_NAME

Requests the user assigned instrument name, useful for distinguishing multiple instruments being controlled by a single client application.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_READ_INSTRUMENT_NAME (0xA6)	1

Response Format (up to 20 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_INSTRUMENT_NAME (0xA7)	1

Byte 1:	Instrument Name (string)	up to 19

## AMS_MSG_READ_BOX_AMOUNT

Requests the serial number of the instrument.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_READ_BOX_AMOUNT(0xA8)	1

Response Format (up to 10 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_BOX_AMOUNT(0xA9)	1

Byte 1:	Amount of Boxes (byte 0 to 8)	1

## AMS_MSG_READ_SERIAL_NUMBER

Requests the serial number of the instrument.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_READ_SERIAL_NUMBER(0xA2)	1

Response Format (up to 10 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_SERIAL_NUMBER(0xA3)	1

Byte 1:	Serial Number (string – up to 8 characters)	up to 9

## AMS_MSG_READ_FLASH_BLOCK

Requests the binary array for all one of the 32  channel Sets. The first eight blocks are for the channels while the ninth block is for the global byte.  Block number BoxNum(0 to 7)or 8 for global . If Block number is MessageTerminator (7F) then all 513 bytes are sent (32*2*8+1)

Data is arranged in 2 bytes per channel (highest byte value is 3F):

Bit	7	6	5	4	3	2	1	0

Ch Data 1st byte=	Empty	Empty	Notch	50Hz	HighP2	HighP1	HighP0	OFF

Ch Data 2nd byte=	Empty	Empty	Gain2	Gain1	Gain0	LowP2	LowP1	LowP0

The last Block has one byte for the global bits

Bit	7	6	5	4	3	2	1	0

Empty	Empty	Empty	Empty	NegBus	CalOn	CalGain1	CalGain0.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_READ_FLASH_ARRAY(0XB1)	1

Byte 1:	Block number	1

Response Format (up to 514 bytes)	Length

Byte 0:	AMS_RESPONSE_READ_FLASH_ARRAY (0xC1)	1

Byte 1-514 	DATA	up to 513

## AMS_MSG_LOAD_FLASH_BLOCK

This will load one of the 32  channel Sets from the flash and set the filter and gain switches as specified by the flash variables. Block number is (0 to 7)or 8. If Block number is MessageTerminator (7F) then all channels are set.

Message Format (1 byte)	Length

Byte 0: 	AMS_MSG_LOAD_FLASH_ARRAY(0XB2)	1

Byte 1:	Block number	1

Response Format (1 byte)	Length

Byte 0:	AMS_RESPONSE_LOAD_FLASH_ARRAY (0xC2)	1

## AMS_MSG_SAVE_FLASH_BLOCK

The computer wants to save current data to flash. This will save one box (or block, or 32 channels) of data at a time. The first eight blocks are for the channels while the ninth block is for the global byte:

Bit	7	6	5	4	3	2	1	0

Ch Data 1st byte=	Empty	Empty	Notch	50Hz	HighP2	HighP1	HighP0	OFF

Ch Data 2nd byte=	Empty	Empty	Gain2	Gain1	Gain0	LowP2	LowP1	LowP0

Global bits

Bit	7	6	5	4	3	2	1	0

Empty	Empty	Empty	Empty	NegBus	CalOn	CalGain1	CalGain0.

The format of the call will be B3 BoxNum(0 to 7)or 8 for global bit then the following 64 bytes will be the channel data

EX:  B3 05 42 00 42 00 42 00 ... 42 00 7F

bytes:  0  1  2    3   4   5   6   7      64 65 66

this example would set the flash block for box 5

Message Format (66 bytes)	Length

Byte 0: 	AMS_MSG_SAVE_FLASH_BLOCK(0XB3)	1

Byte 1:	Block number	1

Byte 2-66	Channel Data	64

Response Format (3 or 66 bytes 3 bytes if global block (block 8))	Length

Byte 0:	AMS_RESPONSE_SAVE_FLASH_BLOCK (0xC3)	1

Byte 1:	Block number	1

Byte 2-66	Channel Data (or byte 2 is just the global byte)	64

## AMS_MSG_WRITE_CHANNEL_VALUES

Modifies a single setting in the currently running program, data is sent in ASCII

Message Format (10 bytes)	Length

Byte 0: 	AMS_MSG_WRITE_CHANNEL_VALUES (0xB5)	1

Byte 1:	channel number high nibble (ASCII 0-F)	1

Byte 2:	channel number low nibble (ASCII 0-F)	1

Byte 3:	Off (0,1) {ASCII 30,31}	1

Byte 4:	High Pass (0,1,2,3,4,5,6,7)	1

Byte 5:	Line Frequency (0,1)	1

Byte 6:	Notch (0,1)	1

Byte 7:	Reference (0,1)	1

Byte 8:	Low Pass (0,1,2,3,4,5,6,7)	1

Byte 9:	Gain (0,1,2,3,4,5,6,7)	1

Response Format (10 bytes)	Length

Byte 0:	AMS_RESPONSE_WRITE_CHANNEL_VALUES (0xC5)	1

Byte 1:	channel number high nibble (ASCII 0-F)	1

Byte 2:	channel number low nibble (ASCII 0-F)	1

Byte 3:	Off (0,1)	1

Byte 4:	High Pass (0,1,2,3,4,5,6,7)	1

Byte 5:	Line Frequency (0,1)	1

Byte 6:	Notch (0,1)	1

Byte 7:	Reference (0,1)	1

Byte 8:	Low Pass (0,1,2,3,4,5,6,7)	1

Byte 9:	Gain (0,1,2,3,4,5,6,7)	1

The following chart lists the settings that correspond to the values transmitted in the channel data block under the standard production configuration of the hardware.  Customized instruments may implement different settings for the data values.  Detailed information regarding specific customized units can be obtained through the {hardware config block}.  The High Pass and Low Pass Filters support 8 settings (data values 0 – 7) and the Gain supports either 11 or 13 settings (data values 0 – 10 or 0 – 12).

Value	High Pass Setting	Low Pass Setting	Gain Setting

0	0.1 Hz	100 Hz	1

1	1 Hz	300 Hz	2

2	3 Hz	500 Hz	5

3	10 Hz	1 kHz	10

4	30 Hz	3 kHz	20

5	100 Hz	5 kHz	50

6	300 Hz	10 kHz	100

7	500 Hz	20 kHz	200

## {hardware config block}

The hardware configuration block provides details regarding any hardware modifications made to the instrument.  This information is sufficient to present a customized user interface that accurately represents the capabilities present in the hardware.

The current value for Layout Revision (Byte 0) is 0x01, which represents the data schema detailed in the following table.  When a value of 0x00 is received for Configuration Code (Byte 1), the standard production values for Filter and Gain options should be used and the content of the remaining bytes in this message are undefined.

Hardware Config Block Format (320 bytes)	Length

Byte 0:	Layout Revision	1

Byte 1:	Configuration Code	1

0 = standard production hardware

1 = custom configuration

Byte 2-65:	Each channel can have one of 4 custom Settings	64

Each byte has 8 bits, two bits for each channel

0=Custom 0     1=Custom 1

2=Custom 2     3=Custom 3

Byte:  msb  ##    ##   ##   ##  lsb

= channel  n+3 n+2 n+1  n

so byte 2 would have channel 0, 1, 2, and 3’s custom setting

Byte 66-119	Reserved	54

Byte 120:	{config value} Calibration Gain Setting 0	2

Byte 122:	{config value} Calibration Gain Setting 1	2

Byte 124:	{config value} Calibration Gain Setting 2	2

Byte 126:	{config value} Calibration Gain Setting 3	2

Byte 128-175 	{channel config block} for Custom 0	48

Byte 176-223	{channel config block} for Custom 1	48

Byte 224-271 	{channel config block} for Custom 2	48

Byte 272-319	{channel config block} for Custom 3	48

## {channel config block}

Channel Config Block Format (48 bytes)	Length

Byte 0:	{config value} High Pass Setting 0	2

…	…	…

Byte 15:	{config value} High Pass Setting 7	2

Byte 16:	{config value} Low Pass Setting 0	2

…	…	…

Byte 31:	{config value} Low Pass Setting 7	2

Byte 32:	{config value} Gain Setting 0	2

…	…	…

Byte 47:	{config value} Gain Setting 7	2

## {config value}

Each config values uses 2 bytes to encode a floating point value.  To construct the final value, use the formula M * 10 ^ E where M is the mantissa encoded in byte 0 and E is the exponent encoded in byte 1.

Byte	Mask	Description	# of Bits

0	0x80	Reserved (must be 0)	1

0	0x7F	mantissa (must be in range 1 – 99)	7

(byte 0 =0 to 0x63)

1	0x80	Reserved (must be 0)	1

1	0x40	1 = negative exponent	1

0 = positive exponent

1	0x3F	exponent ( 0 to 15)	6

(byte1= 0 to 0x4F)

# Revision History

  - First Revision